## Cover
Remediation Plan - CAPA For Compliance Assessment at CSL Behring, Broadmeadows, Australia

## CAPAs
Site | SystemID | FID | OLD FID | System # | CSL ID | Finding ID | Finding Description | Priority ID | Priority | CAPA ID | Modified CAPA | CAPA Description
BRN | BRN-LAS_051 | BRN-LAS_051-I1-1-5 | #N/A | LAS_051 | CSL ID | I1 | The process output (result) is being stored in the SW only temporarily. After pre-determined time, it is being deleted. The data is manually transferred by the operator to NuGenesis after each measurement. No info about Methods backup | 1 | HIGH | 5 | 5 | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_052 | BRN-LAS_052-S1-1-1 | #N/A | LAS_052 | 33557 | S1 | System does not allow to configure unique/personal accounts | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_052 | BRN-LAS_052-S2-2-2 | #N/A | LAS_052 | 33557 | S2 | System does not allow to configure multiple access levels (e.g. User Profile) | 2 | SIGNIFICANT | 2 |  | Investigate with the supplier if it's possible to upgrade the system's version and configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_052 | BRN-LAS_052-S3-3-3 | #N/A | LAS_052 | 33557 | S3 | System does not allow to configure any Security feature | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_052 | BRN-LAS_052-S4-2-4 | #N/A | LAS_052 | 33557 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_052 | BRN-LAS_052-T1-1-6 | #N/A | LAS_052 | 33557 | T1 | Audit Trail is not accessible or not configured within the system | 1 | HIGH | 6 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_052 | BRN-LAS_052-T3-2-7 | #N/A | LAS_052 | 33557 | T3 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 7 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_052 | BRN-LAS_052-T4-2-8 | #N/A | LAS_052 | 33557 | T4 | Original Raw Data is not maintained in the SW due to the system's limitation. However, the data is being transferred by the lab operator to the NuGenesis platform | 2 | SIGNIFICANT | 8 |  | Implement a process of automatic data transfer following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual data transfer
BRN | BRN-LAS_052 | BRN-LAS_052-VD1-1-9 | #N/A | LAS_052 | 33557 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_052 | BRN-LAS_052-I1-1-5 | #N/A | LAS_052 | 33557 | I1 | The process output (result) is being stored in the SW only temporarily. After pre-determined time, it is being deleted. The data is manually transferred by the operator to NuGenesis after each measurement. No info about Methods backup | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_053 | BRN-LAS_053-S1-1-1 | #N/A | LAS_053 | 30747 | S1 | System does not allow to configure unique/personal accounts | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_053 | BRN-LAS_053-S2-2-2 | #N/A | LAS_053 | 30747 | S2 | System does not allow to configure multiple access levels (e.g. User Profile) | 2 | SIGNIFICANT | 2 |  | Investigate with the supplier if it's possible to upgrade the system's version and configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_053 | BRN-LAS_053-S3-3-3 | #N/A | LAS_053 | 30747 | S3 | System does not allow to configure any Security feature | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_053 | BRN-LAS_053-S4-2-4 | #N/A | LAS_053 | 30747 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_053 | BRN-LAS_053-T1-1-6 | #N/A | LAS_053 | 30747 | T1 | Audit Trail is not accessible or not configured within the system | 1 | HIGH | 6 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_053 | BRN-LAS_053-T3-2-7 | #N/A | LAS_053 | 30747 | T3 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 7 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_053 | BRN-LAS_053-T4-2-8 | #N/A | LAS_053 | 30747 | T4 | Original Raw Data is not maintained in the SW due to the system's limitation. However, the data is being transferred by the lab operator to the NuGenesis platform | 2 | SIGNIFICANT | 8 |  | Implement a process of automatic data transfer following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual data transfer
BRN | BRN-LAS_053 | BRN-LAS_053-VD1-1-9 | #N/A | LAS_053 | 30747 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_053 | BRN-LAS_053-I1-1-5 | #N/A | LAS_053 | 30747 | I1 | The process output (result) is being stored in the SW only temporarily. After pre-determined time, it is being deleted. The data is manually transferred by the operator to NuGenesis after each measurement. No info about Methods backup | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_054 | BRN-LAS_054-S1-1-1 | #N/A | LAS_054 | 30317 | S1 | System does not allow to configure unique/personal accounts | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_054 | BRN-LAS_054-S2-2-2 | #N/A | LAS_054 | 30317 | S2 | System does not allow to configure multiple access levels (e.g. User Profile) | 2 | SIGNIFICANT | 2 |  | Investigate with the supplier if it's possible to upgrade the system's version and configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_054 | BRN-LAS_054-S3-3-3 | #N/A | LAS_054 | 30317 | S3 | System does not allow to configure any Security feature | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_054 | BRN-LAS_054-S4-2-4 | #N/A | LAS_054 | 30317 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_054 | BRN-LAS_054-T1-1-6 | #N/A | LAS_054 | 30317 | T1 | Audit Trail is not accessible or not configured within the system | 1 | HIGH | 6 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_054 | BRN-LAS_054-T3-2-7 | #N/A | LAS_054 | 30317 | T3 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 7 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_054 | BRN-LAS_054-T4-2-8 | #N/A | LAS_054 | 30317 | T4 | Original Raw Data is not maintained in the SW due to the system's limitation. However, the data is being transferred by the lab operator to the NuGenesis platform | 2 | SIGNIFICANT | 8 |  | Implement a process of automatic data transfer following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual data transfer
BRN | BRN-LAS_054 | BRN-LAS_054-VD1-1-9 | #N/A | LAS_054 | 30317 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_054 | BRN-LAS_054-I1-1-5 | #N/A | LAS_054 | 30317 | I1 | Backup through NuGenesis system every 48 hours. | 1 | HIGH | 5 |  | Ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_054 | BRN-LAS_054-I2-1-6 | #N/A | LAS_054 | 30317 | I2 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_054 | BRN-LAS_054-I3-1-7 | #N/A | LAS_054 | 30317 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_058 | BRN-LAS_058-S1-1-1 | BRN-LAS_058-S1-2-1 | LAS_058 | 29363 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 | 1 | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_058 | BRN-LAS_058-S2-2-2 | #N/A | LAS_058 | 29363 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 | 5 | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_058 | BRN-LAS_058-S3-3-3 | #N/A | LAS_058 | 29363 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 | 6 | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_058 | BRN-LAS_058-S4-2-4 | BRN-LAS_058-S4-3-2 | LAS_058 | 29363 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 | 2 | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_058 | BRN-LAS_058-T1-2-8 | BRN-LAS_058-T1-2-4 | LAS_058 | 29363 | T1 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 8 | 4 | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_058 | BRN-LAS_058-T3-3-9 | #N/A | LAS_058 | 29363 | T3 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 9 | 7 | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_058 | BRN-LAS_058-VD1-1-10 | #N/A | LAS_058 | 29363 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 | 8 | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_058 | BRN-LAS_058-I1-1-5 | #N/A | LAS_058 | 29363 | I1 | Backup through NuGenesis system every 48 hours. | 1 | HIGH | 5 | 9 | Ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_058 | BRN-LAS_058-I2-1-6 | #N/A | LAS_058 | 29363 | I2 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 6 | 10 | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_058 | BRN-LAS_058-I3-1-7 | BRN-LAS_058-I3-2-3 | LAS_058 | 29363 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 7 | 3 | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_060 | BRN-LAS_060-S1-1-1 | #N/A | LAS_060 | 28298 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_060 | BRN-LAS_060-S2-2-2 | #N/A | LAS_060 | 28298 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_060 | BRN-LAS_060-S3-3-3 | #N/A | LAS_060 | 28298 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_060 | BRN-LAS_060-S4-2-4 | #N/A | LAS_060 | 28298 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_060 | BRN-LAS_060-T1-1-8 | #N/A | LAS_060 | 28298 | T1 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 1 | HIGH | 8 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_060 | BRN-LAS_060-T3-2-9 | #N/A | LAS_060 | 28298 | T3 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_060 | BRN-LAS_060-VD1-1-10 | #N/A | LAS_060 | 28298 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_060 | BRN-LAS_060-I1-1-5 | #N/A | LAS_060 | 28298 | I1 | Backup through NuGenesis system every 48 hours. | 1 | HIGH | 5 |  | Ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_060 | BRN-LAS_060-I2-1-6 | #N/A | LAS_060 | 28298 | I2 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_060 | BRN-LAS_060-I3-1-7 | #N/A | LAS_060 | 28298 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_061 | BRN-LAS_061-S1-1-1 | #N/A | LAS_061 | 30340 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_061 | BRN-LAS_061-S2-2-2 | #N/A | LAS_061 | 30340 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_061 | BRN-LAS_061-S3-3-3 | #N/A | LAS_061 | 30340 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_061 | BRN-LAS_061-S4-2-4 | BRN-LAS_061-S4-4-1 | LAS_061 | 30340 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_061 | BRN-LAS_061-T1-2-8 | #N/A | LAS_061 | 30340 | T1 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 8 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_061 | BRN-LAS_061-T3-3-9 | BRN-LAS_061-T3-4-2 | LAS_061 | 30340 | T3 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_061 | BRN-LAS_061-VD1-1-10 | #N/A | LAS_061 | 30340 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_061 | BRN-LAS_061-I1-1-5 | #N/A | LAS_061 | 30340 | I1 | Backup through NuGenesis system every 48 hours. | 1 | HIGH | 5 |  | Ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_061 | BRN-LAS_061-I2-1-6 | #N/A | LAS_061 | 30340 | I2 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_061 | BRN-LAS_061-I3-1-7 | #N/A | LAS_061 | 30340 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_063 | BRN-LAS_063-S1-1-1 | BRN-LAS_063-S1-2-1 | LAS_063 | 30797 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_063 | BRN-LAS_063-S2-2-2 | #N/A | LAS_063 | 30797 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_063 | BRN-LAS_063-S3-3-3 | #N/A | LAS_063 | 30797 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_063 | BRN-LAS_063-S4-2-4 | BRN-LAS_063-S4-3-2 | LAS_063 | 30797 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_063 | BRN-LAS_063-T1-2-8 | BRN-LAS_063-T1-2-4 | LAS_063 | 30797 | T1 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 8 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_063 | BRN-LAS_063-T3-3-9 | #N/A | LAS_063 | 30797 | T3 | Present a log of changes not complete. Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_063 | BRN-LAS_063-VD1-1-10 | #N/A | LAS_063 | 30797 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_063 | BRN-LAS_063-I2-1-5 | #N/A | LAS_063 | 30797 | I2 | Data not protected | 1 | HIGH | 5 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_063 | BRN-LAS_063-I3-1-6 | #N/A | LAS_063 | 30797 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_069 | BRN-LAS_069-S1-1-1 | #N/A | LAS_069 | 30707 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_069 | BRN-LAS_069-S2-2-2 | #N/A | LAS_069 | 30707 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_069 | BRN-LAS_069-S3-3-3 | #N/A | LAS_069 | 30707 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_069 | BRN-LAS_069-S4-2-4 | BRN-LAS_069-S4-2-1 | LAS_069 | 30707 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_069 | BRN-LAS_069-T1-1-7 | #N/A | LAS_069 | 30707 | T1 | It has been not possible to check if the methods can have the same name or there is a versioning | 1 | HIGH | 7 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_069 | BRN-LAS_069-T2-1-8 | #N/A | LAS_069 | 30707 | T2 | Time Zone not protected | 1 | HIGH | 8 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_069 | BRN-LAS_069-T3-3-9 | BRN-LAS_069-T3-2-2 | LAS_069 | 30707 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_069 | BRN-LAS_069-VD1-1-10 | #N/A | LAS_069 | 30707 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_069 | BRN-LAS_069-I1-2-3 | #N/A | LAS_069 | 30707 | I1 | Automatic Backup performed one time every six months. | 2 | SIGNIFICANT | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_074 | BRN-LAS_074-S3-4-1 | #N/A | LAS_074 | 34154 | S3 | Security features not compliant | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_074 | BRN-LAS_074-S4-3-2 | BRN-LAS_074-S4-4-1 | LAS_074 | 34154 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_074 | BRN-LAS_074-T2-2-4 | #N/A | LAS_074 | 34154 | T2 | Time reference (Time zone) is not protected. Date and Time protected. | 2 | SIGNIFICANT | 4 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_074 | BRN-LAS_074-T3-3-5 | BRN-LAS_074-T3-4-2 | LAS_074 | 34154 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 5 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_074 | BRN-LAS_074-VD1-2-6 | #N/A | LAS_074 | 34154 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 6 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_076 | BRN-LAS_076-S3-4-1 | #N/A | LAS_076 | 33526 | S3 | Security features not compliant | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_076 | BRN-LAS_076-S4-3-2 | BRN-LAS_076-S4-3-1 | LAS_076 | 33526 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_076 | BRN-LAS_076-T3-3-3 | BRN-LAS_076-T3-3-2 | LAS_076 | 33526 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 3 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_076 | BRN-LAS_076-VD1-2-4 | #N/A | LAS_076 | 33526 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 4 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_089 | BRN-LAS_089-I1-1-1 | #N/A | LAS_089 | 20008692 | I1 | Automatic backup not accessible or not configured. | 1 | HIGH | 1 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_089 | BRN-LAS_089-I3-1-2 | #N/A | LAS_089 | 20008692 | I3 | Data can be modified within OS | 1 | HIGH | 2 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_089 | BRN-LAS_089-T3-2-3 | BRN-LAS_089-T3-2-2 | LAS_089 | 20008692 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 3 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_089 | BRN-LAS_089-VD1-1-4 | #N/A | LAS_089 | 20008692 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 4 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_090 | BRN-LAS_090-I1-2-1 | #N/A | LAS_090 | 20050245 | I1 | Automatic backup not accessible or not configured. | 2 | SIGNIFICANT | 1 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_090 | BRN-LAS_090-I3-2-2 | #N/A | LAS_090 | 20050245 | I3 | Data can be modified within OS | 2 | SIGNIFICANT | 2 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_090 | BRN-LAS_090-T3-3-3 | BRN-LAS_090-T3-2-2 | LAS_090 | 20050245 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 3 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_090 | BRN-LAS_090-VD1-2-4 | #N/A | LAS_090 | 20050245 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 4 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_090 | BRN-LAS_090-I1-1-5 | #N/A | LAS_090 | 20050245 | I1 | Manual backup performed one time every two months. | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_090 | BRN-LAS_090-I2-1-6 | #N/A | LAS_090 | 20050245 | I2 | Data can be deleted by the SW application | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_090 | BRN-LAS_090-I3-1-7 | #N/A | LAS_090 | 20050245 | I3 | Data can be deleted within OS but the SOP says so to copy. For output record mitigation action in place to copy all the relevant GxP information into NuGenesis files repository | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_097 | BRN-LAS_097-S1-1-1 | #N/A | LAS_097 | 29679 | S1 | SW personal account. Administrator generic account present. One profile has double account: Administrator and Operator | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_097 | BRN-LAS_097-S2-2-2 | #N/A | LAS_097 | 29679 | S2 | Administrator generic account present. One profile has double account: Administrator and Operator. Segregation of Duties not assured due to system limitation | 2 | SIGNIFICANT | 2 |  | Investigate with the supplier if it's possible to upgrade the system's version and configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_097 | BRN-LAS_097-S3-3-3 | #N/A | LAS_097 | 29679 | S3 | System does not allow to configure any Security feature | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_097 | BRN-LAS_097-S4-2-4 | #N/A | LAS_097 | 29679 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_097 | BRN-LAS_097-T1-1-8 | #N/A | LAS_097 | 29679 | T1 | No automatic Audit Trail or automatic data versioning is allowed by the SW application. Data can be modified/overwritten without any automatic traceability | 1 | HIGH | 8 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_097 | BRN-LAS_097-T3-2-9 | #N/A | LAS_097 | 29679 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_097 | BRN-LAS_097-VD1-1-10 | #N/A | LAS_097 | 29679 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_097 | BRN-LAS_097-I1-2-5 | #N/A | LAS_097 | 29679 | I1 | Manual backup performed one time every six months. | 2 | SIGNIFICANT | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_097 | BRN-LAS_097-I2-2-6 | #N/A | LAS_097 | 29679 | I2 | Data can be deleted by the SW application | 2 | SIGNIFICANT | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_097 | BRN-LAS_097-I3-2-7 | #N/A | LAS_097 | 29679 | I3 | Data can be deleted within OS but the SOP says so to copy. For output record mitigation action in place to copy all the relevant GxP information into NuGenesis files repository | 2 | SIGNIFICANT | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_098 | BRN-LAS_098-S1-2-1 | #N/A | LAS_098 | 32484 | S1 | SW personal account. Administrator generic account present. | 2 | SIGNIFICANT | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_098 | BRN-LAS_098-S2-3-2 | #N/A | LAS_098 | 32484 | S2 | Administrator generic account present. One profile has double account: Administrator and Operator. Segregation of Duties not assured due to system limitation | 3 | MODERATE | 2 |  | Investigate with the supplier if it's possible to upgrade the system's version and configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_098 | BRN-LAS_098-S3-4-3 | #N/A | LAS_098 | 32484 | S3 | System does not allow to configure any Security feature | 4 | LOW | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_098 | BRN-LAS_098-S4-3-4 | #N/A | LAS_098 | 32484 | S4 | No Official User List is in place | 3 | MODERATE | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_098 | BRN-LAS_098-T1-2-8 | #N/A | LAS_098 | 32484 | T1 | No automatic Audit Trail or automatic data versioning is allowed by the SW application. Data can be modified/overwritten without any automatic traceability | 2 | SIGNIFICANT | 8 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_098 | BRN-LAS_098-T3-3-9 | #N/A | LAS_098 | 32484 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_098 | BRN-LAS_098-VD1-2-10 | #N/A | LAS_098 | 32484 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_098 | BRN-LAS_098-I1-2-3 | #N/A | LAS_098 | 32484 | I1 | Automatic backup not in place | 2 | SIGNIFICANT | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_098 | BRN-LAS_098-I3-2-4 | #N/A | LAS_098 | 32484 | I3 | Data can be altered within the OS. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_100 | BRN-LAS_100-S1-2-1 | #N/A | LAS_100 | 33692 | S1 | System is limited to 20 users. Generic access for admin level. | 2 | SIGNIFICANT | 1 |  | If the generic account is necessary for business reasons, document and map its functionality and access capabilities, and grant traceability of its actions
BRN | BRN-LAS_100 | BRN-LAS_100-S3-4-2 | #N/A | LAS_100 | 33692 | S3 | Security features not compliant | 4 | LOW | 2 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_100 | BRN-LAS_100-T1-2-5 | #N/A | LAS_100 | 33692 | T1 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 5 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_100 | BRN-LAS_100-T2-2-6 | #N/A | LAS_100 | 33692 | T2 | Time reference is not protected on the PC, HMI time reference protected. | 2 | SIGNIFICANT | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_100 | BRN-LAS_100-T3-3-7 | BRN-LAS_100-T3-4-2 | LAS_100 | 33692 | T3 | Audit Trail is not accessible or not configured within the system | 3 | MODERATE | 7 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_100 | BRN-LAS_100-VD1-2-8 | #N/A | LAS_100 | 33692 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_100 | BRN-LAS_100-I2-2-3 | #N/A | LAS_100 | 33692 | I2 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_100 | BRN-LAS_100-I3-2-4 | #N/A | LAS_100 | 33692 | I3 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_105 | BRN-LAS_105-S3-4-1 | #N/A | LAS_105 | 30536 | S3 | Security features not compliant | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_105 | BRN-LAS_105-S4-3-2 | #N/A | LAS_105 | 30536 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_105 | BRN-LAS_105-T1-2-5 | #N/A | LAS_105 | 30536 | T1 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 5 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_105 | BRN-LAS_105-T3-3-6 | #N/A | LAS_105 | 30536 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 6 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_105 | BRN-LAS_105-VD1-2-7 | #N/A | LAS_105 | 30536 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 7 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_105 | BRN-LAS_105-I2-2-3 | #N/A | LAS_105 | 30536 | I2 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_105 | BRN-LAS_105-I3-2-4 | #N/A | LAS_105 | 30536 | I3 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_106 | BRN-LAS_106-S3-4-1 | #N/A | LAS_106 | 30925 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_106 | BRN-LAS_106-S4-3-2 | #N/A | LAS_106 | 30925 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_106 | BRN-LAS_106-T1-2-5 | #N/A | LAS_106 | 30925 | T1 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 5 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_106 | BRN-LAS_106-T3-3-6 | #N/A | LAS_106 | 30925 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 6 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_106 | BRN-LAS_106-VD1-2-7 | #N/A | LAS_106 | 30925 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 7 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_114 | BRN-LAS_114-S1-1-1 | #N/A | LAS_114 | 20168054 | S1 | Generic account configured in the system | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_114 | BRN-LAS_114-S2-2-2 | #N/A | LAS_114 | 20168054 | S2 | There are two different type of administrator. To be verified if one of these is involved in the process | 2 | SIGNIFICANT | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_114 | BRN-LAS_114-S4-2-3 | #N/A | LAS_114 | 20168054 | S4 | Users reported that there is a list but should be updated | 2 | SIGNIFICANT | 3 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_114 | BRN-LAS_114-T1-2-4 | #N/A | LAS_114 | 20168054 | T1 | User creation is traced. The modification of Method is traced in the logbook BRN-WOI-0002157-F01_2.0 (old and new value) | 2 | SIGNIFICANT | 4 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_114 | BRN-LAS_114-T3-2-5 | #N/A | LAS_114 | 20168054 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 5 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_114 | BRN-LAS_114-VD1-1-6 | #N/A | LAS_114 | 20168054 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 6 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_114 | BRN-LAS_114-I1-3-3 | #N/A | LAS_114 | 20168054 | I1 | Automatic Backup performed one time every six months. | 3 | MODERATE | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_114 | BRN-LAS_114-I2-3-4 | #N/A | LAS_114 | 20168054 | I2 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_114 | BRN-LAS_114-I3-3-5 | #N/A | LAS_114 | 20168054 | I3 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_114 | BRN-LAS_114-I4-3-6 | #N/A | LAS_114 | 20168054 | I4 | No process output. Report is not being generated as a GxP electronic record | 3 | MODERATE | 6 |  | Check the system's behaviour towards all electronic records. If autosave is not enforced by the system, document the manual workaround for saving output records
BRN | BRN-LAS_124 | BRN-LAS_124-S3-5-1 | #N/A | LAS_124 | 30548 | S3 | Security features not compliant | 5 | MINIMUM | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_124 | BRN-LAS_124-S4-4-2 | BRN-LAS_124-S4-4-1 | LAS_124 | 30548 | S4 | No Official User List is in place | 4 | LOW | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_124 | BRN-LAS_124-T3-4-7 | BRN-LAS_124-T3-4-2 | LAS_124 | 30548 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 4 | LOW | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_124 | BRN-LAS_124-VD1-3-8 | #N/A | LAS_124 | 30548 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_124 | BRN-LAS_124-I1-3-3 | #N/A | LAS_124 | 30548 | I1 | Automatic Backup performed one time every six months. | 3 | MODERATE | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_124 | BRN-LAS_124-I2-3-4 | #N/A | LAS_124 | 30548 | I2 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_124 | BRN-LAS_124-I3-3-5 | #N/A | LAS_124 | 30548 | I3 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_124 | BRN-LAS_124-I4-3-6 | #N/A | LAS_124 | 30548 | I4 | No process output. Report is not being generated as a GxP electronic record | 3 | MODERATE | 6 |  | Check the system's behaviour towards all electronic records. If autosave is not enforced by the system, document the manual workaround for saving output records
BRN | BRN-LAS_127 | BRN-LAS_127-S3-5-1 | #N/A | LAS_127 | 33320 | S3 | Security features not compliant | 5 | MINIMUM | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_127 | BRN-LAS_127-S4-4-2 | BRN-LAS_127-S4-4-1 | LAS_127 | 33320 | S4 | No Official User List is in place | 4 | LOW | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_127 | BRN-LAS_127-T3-4-7 | BRN-LAS_127-T3-4-2 | LAS_127 | 33320 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 4 | LOW | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_127 | BRN-LAS_127-VD1-3-8 | #N/A | LAS_127 | 33320 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_127 | BRN-LAS_127-I1-3-3 | #N/A | LAS_127 | 33320 | I1 | Automatic Backup performed one time every six months. | 3 | MODERATE | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_127 | BRN-LAS_127-I2-3-4 | #N/A | LAS_127 | 33320 | I2 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_127 | BRN-LAS_127-I3-3-5 | #N/A | LAS_127 | 33320 | I3 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_127 | BRN-LAS_127-I4-3-6 | #N/A | LAS_127 | 33320 | I4 | No process output. Report is not being generated as a GxP electronic record | 3 | MODERATE | 6 |  | Check the system's behaviour towards all electronic records. If autosave is not enforced by the system, document the manual workaround for saving output records
BRN | BRN-LAS_130 | BRN-LAS_130-S3-5-1 | #N/A | LAS_130 | 20134451 | S3 | Security features not compliant | 5 | MINIMUM | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_130 | BRN-LAS_130-S4-4-2 | BRN-LAS_130-S4-4-1 | LAS_130 | 20134451 | S4 | No Official User List is in place | 4 | LOW | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_130 | BRN-LAS_130-T3-4-7 | BRN-LAS_130-T3-4-2 | LAS_130 | 20134451 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 4 | LOW | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_130 | BRN-LAS_130-VD1-3-8 | #N/A | LAS_130 | 20134451 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_130 | BRN-LAS_130-I1-3-3 | #N/A | LAS_130 | 20134451 | I1 | Automatic Backup performed one time every six months. | 3 | MODERATE | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_130 | BRN-LAS_130-I2-3-4 | #N/A | LAS_130 | 20134451 | I2 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_130 | BRN-LAS_130-I3-3-5 | #N/A | LAS_130 | 20134451 | I3 | No process output. Report is not being generated as a GxP electronic record. No info about how input record is managed. | 3 | MODERATE | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_130 | BRN-LAS_130-I4-3-6 | #N/A | LAS_130 | 20134451 | I4 | No process output. Report is not being generated as a GxP electronic record | 3 | MODERATE | 6 |  | Check the system's behaviour towards all electronic records. If autosave is not enforced by the system, document the manual workaround for saving output records
BRN | BRN-LAS_131 | BRN-LAS_131-S3-5-1 | #N/A | LAS_131 | 20134450 | S3 | Security features not compliant | 5 | MINIMUM | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_131 | BRN-LAS_131-S4-4-2 | BRN-LAS_131-S4-4-1 | LAS_131 | 20134450 | S4 | No Official User List is in place | 4 | LOW | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_131 | BRN-LAS_131-T3-4-7 | BRN-LAS_131-T3-4-2 | LAS_131 | 20134450 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 4 | LOW | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_131 | BRN-LAS_131-VD1-3-8 | #N/A | LAS_131 | 20134450 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_131 | BRN-LAS_131-I1-3-5 | #N/A | LAS_131 | 20134450 | I1 | No automatic backup in place (The user is trained to put the relevant gxp data into specific folder that in regularly backed up) | 3 | MODERATE | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_131 | BRN-LAS_131-I3-3-6 | #N/A | LAS_131 | 20134450 | I3 | Data can be altered within the OS. | 3 | MODERATE | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_145 | BRN-LAS_145-S1-3-1 | #N/A | LAS_145 | 33328 | S1 | Generic login for SW application | 3 | MODERATE | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_145 | BRN-LAS_145-S2-4-2 | #N/A | LAS_145 | 33328 | S2 | 3 groups configured in the system, however the access to SW is generic so the proper segregation of duties cannot be assured | 4 | LOW | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_145 | BRN-LAS_145-S3-5-3 | #N/A | LAS_145 | 33328 | S3 | System does not allow to configure any Security feature | 5 | MINIMUM | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_145 | BRN-LAS_145-S4-4-4 | #N/A | LAS_145 | 33328 | S4 | No Official User List is in place | 4 | LOW | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_145 | BRN-LAS_145-T1-3-7 | #N/A | LAS_145 | 33328 | T1 | Audit Trail is not accessible or not configured within the system | 3 | MODERATE | 7 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_145 | BRN-LAS_145-T2-3-8 | #N/A | LAS_145 | 33328 | T2 | Time reference (Time zone) is not protected. Date and Time protected. | 3 | MODERATE | 8 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_145 | BRN-LAS_145-T3-4-9 | #N/A | LAS_145 | 33328 | T3 | Audit Trail is not accessible or not configured within the system | 4 | LOW | 9 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_145 | BRN-LAS_145-VD1-3-10 | #N/A | LAS_145 | 33328 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_145 | BRN-LAS_145-ES1-1-6 | #N/A | LAS_145 | 33328 | ES1 | Information not provided | 1 | HIGH | 6 |  | Verify the ES review procedure, if present. If it's not, implement a process to periodically review the correct execution of Electronic Signatures
BRN | BRN-LAS_145 | BRN-LAS_145-ES2-1-7 | #N/A | LAS_145 | 33328 | ES2 | Information not provided | 1 | HIGH | 7 |  | Verify the electronically signed reports procedure, if present, and ensure that there is a documented periodic procedure to check that they are signed by the user that created them and their creation is controlled
BRN | BRN-LAS_145 | BRN-LAS_145-ES3-1-8 | #N/A | LAS_145 | 33328 | ES3 | Information not provided | 1 | HIGH | 8 |  | Verify that reports can be configured to include ES meaning
BRN | BRN-LAS_145 | BRN-LAS_145-I2-1-4 | #N/A | LAS_145 | 33328 | I2 | Information not provided | 1 | HIGH | 4 |  | Verify the status of the system's compliance. If necessary, change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_168 | BRN-LAS_168-S1-1-1 | #N/A | LAS_168 | 34431 | S1 | Under Active Directory (using NuGenesis) Multiple accounts set for same users: 1st case) Double account: MT Service and Administrator 2nd case) Double account: Admin + Active Directory gesteuert | 1 | HIGH | 1 |  | If the generic account is necessary for business reasons, document and map its functionality and access capabilities, and grant traceability of its actions
BRN | BRN-LAS_168 | BRN-LAS_168-S2-2-2 | #N/A | LAS_168 | 34431 | S2 | Active directory Multiple accounts set for same users: 1st case) Double account: MT Service and Administrator 2nd case) Double account: Admin + Active Directory gesteuert | 2 | SIGNIFICANT | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_168 | BRN-LAS_168-S4-2-3 | BRN-LAS_168-S4-2-1 | LAS_168 | 34431 | S4 | Every year is done the access periodic review | 2 | SIGNIFICANT | 3 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_168 | BRN-LAS_168-T3-2-5 | BRN-LAS_168-T3-2-2 | LAS_168 | 34431 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 5 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_168 | BRN-LAS_168-VD1-1-9 | #N/A | LAS_168 | 34431 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_168 | BRN-LAS_168-I2-2-3 | #N/A | LAS_168 | 34431 | I2 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_168 | BRN-LAS_168-I3-2-4 | #N/A | LAS_168 | 34431 | I3 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_201 | BRN-LAS_201-S3-4-1 | #N/A | LAS_201 | 31899 | S3 | Security features not compliant | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_201 | BRN-LAS_201-S4-3-2 | #N/A | LAS_201 | 31899 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_201 | BRN-LAS_201-T3-3-5 | #N/A | LAS_201 | 31899 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 5 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_201 | BRN-LAS_201-VD1-2-6 | #N/A | LAS_201 | 31899 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 6 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_201 | BRN-LAS_201-I2-2-9 | #N/A | LAS_201 | 31899 | I2 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 9 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_201 | BRN-LAS_201-I3-2-10 | #N/A | LAS_201 | 31899 | I3 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 10 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_206 | BRN-LAS_206-S3-4-7 | #N/A | LAS_206 | 20134448 | S3 | Security features not compliant | 4 | LOW | 7 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_206 | BRN-LAS_206-S4-3-8 | #N/A | LAS_206 | 20134448 | S4 | No Official User List is in place | 3 | MODERATE | 8 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_206 | BRN-LAS_206-T3-3-11 | #N/A | LAS_206 | 20134448 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 11 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_206 | BRN-LAS_206-VD1-2-12 | #N/A | LAS_206 | 20134448 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 12 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_206 | BRN-LAS_206-I2-2-3 | #N/A | LAS_206 | 20134448 | I2 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_206 | BRN-LAS_206-I3-2-4 | #N/A | LAS_206 | 20134448 | I3 | Data is automatically transferred to NuGenesis. On the local PC, there is only a copy of the data. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_207 | BRN-LAS_207-S3-4-1 | #N/A | LAS_207 | 20134449 | S3 | Security features not compliant | 4 | LOW | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_207 | BRN-LAS_207-S4-3-2 | #N/A | LAS_207 | 20134449 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_207 | BRN-LAS_207-T3-3-5 | #N/A | LAS_207 | 20134449 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 5 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_207 | BRN-LAS_207-VD1-2-6 | #N/A | LAS_207 | 20134449 | VD1 | Computer System Validation executed but incomplete | 2 | HIGH | 6 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_207 | BRN-LAS_207-I3-3-5 | #N/A | LAS_207 | 20134449 | I3 | User can delete the data on the OS level | 3 | MODERATE | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_208 | BRN-LAS_208-S1-3-1 | #N/A | LAS_208 | 31850 | S1 | User management settings are not accessible or the most privileged access available at the time of the assessment has no access. | 3 | MODERATE | 1 |  | Verify the system's access capability and user management compliance. If it's not, implement a proper user management functionality with individual and personal accounts
BRN | BRN-LAS_208 | BRN-LAS_208-S2-4-2 | #N/A | LAS_208 | 31850 | S2 | User management settings are not accessible or the most privileged access available at the time of the assessment has no access. | 4 | LOW | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_208 | BRN-LAS_208-S3-5-3 | #N/A | LAS_208 | 31850 | S3 | Security features are not accessible or the most privileged access available at the time of the assessment has no access. | 5 | MINIMUM | 3 |  | Verify the system's security features compliance. If it's not, configure security features according to relevant SOP and good practices
BRN | BRN-LAS_208 | BRN-LAS_208-S4-4-4 | BRN-LAS_208-S4-3-1 | LAS_208 | 31850 | S4 | No Official User List is in place | 4 | LOW | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_208 | BRN-LAS_208-T1-3-6 | #N/A | LAS_208 | 31850 | T1 | Audit Trail either not accessible or not configured. | 3 | MODERATE | 6 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_208 | BRN-LAS_208-T2-3-7 | #N/A | LAS_208 | 31850 | T2 | Time zone can be changed | 3 | MODERATE | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_208 | BRN-LAS_208-T3-4-8 | BRN-LAS_208-T3-3-2 | LAS_208 | 31850 | T3 | Audit Trail either not accessible or not configured. | 4 | LOW | 8 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_208 | BRN-LAS_208-VD1-3-9 | #N/A | LAS_208 | 31850 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_208 | BRN-LAS_208-I3-2-5 | #N/A | LAS_208 | 31850 | I3 | Data Folder is not protected from alteration/deletion within OS. | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_216 | BRN-LAS_216-S1-2-1 | #N/A | LAS_216 | 29930 | S1 | Users are using generic account (Operator, Supervisor, Administrator) | 2 | SIGNIFICANT | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_216 | BRN-LAS_216-S2-3-2 | #N/A | LAS_216 | 29930 | S2 | 3 groups configured in the system, however the access to SW is generic so the proper segregation of duties cannot be assured | 3 | MODERATE | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_216 | BRN-LAS_216-S3-4-3 | #N/A | LAS_216 | 29930 | S3 | Security features not compliant | 4 | LOW | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_216 | BRN-LAS_216-S4-3-4 | BRN-LAS_216-S4-3-1 | LAS_216 | 29930 | S4 | No Official User List is in place | 3 | MODERATE | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_216 | BRN-LAS_216-T1-2-6 | #N/A | LAS_216 | 29930 | T1 | User is generic, no traceability to a specific user in generated Audit Trail. Audit Trail is protected from any alternation. | 2 | SIGNIFICANT | 6 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_216 | BRN-LAS_216-T2-2-7 | #N/A | LAS_216 | 29930 | T2 | Time reference is not protected. | 2 | SIGNIFICANT | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_216 | BRN-LAS_216-T3-3-8 | BRN-LAS_216-T3-3-2 | LAS_216 | 29930 | T3 | Audit Trail is not periodically reviewed. | 3 | MODERATE | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_216 | BRN-LAS_216-VD1-2-9 | #N/A | LAS_216 | 29930 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_216 | BRN-LAS_216-I3-2-5 | #N/A | LAS_216 | 29930 | I3 | Data Folder is not protected from alteration/deletion within OS. | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_217 | BRN-LAS_217-S1-2-1 | #N/A | LAS_217 | 29931 | S1 | Users are using generic account (Operator, Supervisor, Administrator) | 2 | SIGNIFICANT | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_217 | BRN-LAS_217-S2-3-2 | #N/A | LAS_217 | 29931 | S2 | 3 groups configured in the system, however the access to SW is generic so the proper segregation of duties cannot be assured | 3 | MODERATE | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_217 | BRN-LAS_217-S3-4-3 | #N/A | LAS_217 | 29931 | S3 | Security features not compliant | 4 | LOW | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_217 | BRN-LAS_217-S4-3-4 | BRN-LAS_217-S4-3-1 | LAS_217 | 29931 | S4 | No Official User List is in place | 3 | MODERATE | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_217 | BRN-LAS_217-T1-2-6 | #N/A | LAS_217 | 29931 | T1 | User is generic, no traceability to a specific user in generated Audit Trail. Audit Trail is protected from any alternation. | 2 | SIGNIFICANT | 6 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_217 | BRN-LAS_217-T2-2-7 | #N/A | LAS_217 | 29931 | T2 | Time reference is not protected. | 2 | SIGNIFICANT | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_217 | BRN-LAS_217-T3-3-8 | BRN-LAS_217-T3-3-2 | LAS_217 | 29931 | T3 | Audit Trail is not periodically reviewed. | 3 | MODERATE | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_217 | BRN-LAS_217-VD1-2-9 | #N/A | LAS_217 | 29931 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_217 | BRN-LAS_217-I3-2-5 | #N/A | LAS_217 | 29931 | I3 | Data Folder is not protected from alteration/deletion within OS. | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_218 | BRN-LAS_218-S1-2-1 | #N/A | LAS_218 | 29932 | S1 | Users are using generic account (Operator, Supervisor, Administrator) | 2 | SIGNIFICANT | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_218 | BRN-LAS_218-S2-3-2 | #N/A | LAS_218 | 29932 | S2 | 3 groups configured in the system, however the access to SW is generic so the proper segregation of duties cannot be assured | 3 | MODERATE | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_218 | BRN-LAS_218-S3-4-3 | #N/A | LAS_218 | 29932 | S3 | Security features not compliant | 4 | LOW | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_218 | BRN-LAS_218-S4-3-4 | BRN-LAS_218-S4-3-1 | LAS_218 | 29932 | S4 | No Official User List is in place | 3 | MODERATE | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_218 | BRN-LAS_218-T1-2-6 | #N/A | LAS_218 | 29932 | T1 | User is generic, no traceability to a specific user in generated Audit Trail. Audit Trail is protected from any alternation. | 2 | SIGNIFICANT | 6 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_218 | BRN-LAS_218-T2-2-7 | #N/A | LAS_218 | 29932 | T2 | Time reference is not protected. | 2 | SIGNIFICANT | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_218 | BRN-LAS_218-T3-3-8 | BRN-LAS_218-T3-3-2 | LAS_218 | 29932 | T3 | Audit Trail is not periodically reviewed. | 3 | MODERATE | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_218 | BRN-LAS_218-VD1-2-9 | #N/A | LAS_218 | 29932 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_218 | BRN-LAS_218-ES1-1-4 | #N/A | LAS_218 | 29932 | ES1 | Information not provided | 1 | HIGH | 4 |  | Verify the ES review procedure, if present. If it's not, implement a process to periodically review the correct execution of Electronic Signatures
BRN | BRN-LAS_218 | BRN-LAS_218-ES2-1-5 | #N/A | LAS_218 | 29932 | ES2 | Information not provided | 1 | HIGH | 5 |  | Verify the electronically signed reports procedure, if present, and ensure that there is a documented periodic procedure to check that they are signed by the user that created them and their creation is controlled
BRN | BRN-LAS_218 | BRN-LAS_218-ES3-1-6 | #N/A | LAS_218 | 29932 | ES3 | Information not provided | 1 | HIGH | 6 |  | Verify that reports can be configured to include ES meaning
BRN | BRN-LAS_218 | BRN-LAS_218-I1-1-2 | #N/A | LAS_218 | 29932 | I1 | Info not provided. It should be evaluated together with LARA system | 1 | HIGH | 2 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_223 | BRN-LAS_223-S1-2-1 | #N/A | LAS_223 | QC-09 | S1 | It is present a generic Super User Admin account that is mandatory and the IT department is in control of it. | 2 | SIGNIFICANT | 1 |  | If the generic account is necessary for business reasons, document and map its functionality and access capabilities, and grant traceability of its actions
BRN | BRN-LAS_223 | BRN-LAS_223-T3-2-3 | BRN-LAS_223-T3-2-2 | LAS_223 | QC-09 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 3 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_223 | BRN-LAS_223-VD1-1-7 | #N/A | LAS_223 | QC-09 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 7 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_223 | BRN-LAS_223-I1-3-3 | #N/A | LAS_223 | QC-09 | I1 | Automatic Backup is not in place | 3 | MODERATE | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_223 | BRN-LAS_223-I3-3-4 | #N/A | LAS_223 | QC-09 | I3 | Data can be deleted within OS | 3 | MODERATE | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_283 | BRN-LAS_283-S3-5-1 | #N/A | LAS_283 | 34081 | S3 | System does not allow to configure any Security feature | 5 | MINIMUM | 1 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_283 | BRN-LAS_283-S4-4-2 | BRN-LAS_283-S4-4-1 | LAS_283 | 34081 | S4 | List of the users is not available | 4 | LOW | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_283 | BRN-LAS_283-T1-3-5 | #N/A | LAS_283 | 34081 | T1 | Due to systems limitation, audit trail is not a part of the software. Logbook is in place to trace the activity. | 3 | MODERATE | 5 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_283 | BRN-LAS_283-T2-3-6 | #N/A | LAS_283 | 34081 | T2 | Time zone can be changed | 3 | MODERATE | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_283 | BRN-LAS_283-T3-4-7 | BRN-LAS_283-T3-4-2 | LAS_283 | 34081 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 4 | LOW | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_283 | BRN-LAS_283-VD1-3-8 | #N/A | LAS_283 | 34081 | VD1 | Computer System Validation executed but incomplete | 3 | MODERATE | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_283 | BRN-LAS_283-I1-2-3 | #N/A | LAS_283 | 34081 | I1 | Automatic backup not in place, however the copy of the data is in the server location (EUCHBRN4884W) which is automatically backed up | 2 | SIGNIFICANT | 3 |  | Setup an automatic backup process. If not possible, ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_283 | BRN-LAS_283-I2-2-4 | #N/A | LAS_283 | 34081 | I2 | Data is not protected from alteration/deletion by the SW application. The data is being manually transferred by the end user into the NuGenesis after each measurement so there is a risk mitigation action in place | 2 | SIGNIFICANT | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_283 | BRN-LAS_283-I3-2-5 | #N/A | LAS_283 | 34081 | I3 | Data Folders are not protected from alteration/deletion within OS. The data is being manually transferred by the end user into the NuGenesis after each measurement so there is a risk mitigation action in place | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_293 | BRN-LAS_293-S3-4-1 | #N/A | LAS_293 | 31620 | S3 | Pre-coded security features. Only configurable parameters are Password Aging and automatic logout | 4 | LOW | 1 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_293 | BRN-LAS_293-S4-3-2 | #N/A | LAS_293 | 31620 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_293 | BRN-LAS_293-T2-2-6 | #N/A | LAS_293 | 31620 | T2 | Date and time can be changed | 2 | SIGNIFICANT | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_293 | BRN-LAS_293-T3-3-7 | #N/A | LAS_293 | 31620 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_293 | BRN-LAS_293-VD1-2-8 | #N/A | LAS_293 | 31620 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_293 | BRN-LAS_293-I1-2-3 | #N/A | LAS_293 | 31620 | I1 | Automatic backup not in place, however the copy of the data is in the server location (EUCHBRN4884W) which is automatically backed up | 2 | SIGNIFICANT | 3 |  | Ensure that the modality and frequency of the backups are defined in an official document with a risk-based approach
BRN | BRN-LAS_293 | BRN-LAS_293-I2-2-4 | #N/A | LAS_293 | 31620 | I2 | Data is not protected from alteration/deletion by the SW application. The data is being manually transferred by the end user into the NuGenesis after each measurement so there is a risk mitigation action in place | 2 | SIGNIFICANT | 4 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_293 | BRN-LAS_293-I3-2-5 | #N/A | LAS_293 | 31620 | I3 | Data Folders are not protected from alteration/deletion within OS. The data is being manually transferred by the end user into the NuGenesis after each measurement so there is a risk mitigation action in place | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_294 | BRN-LAS_294-S3-4-1 | #N/A | LAS_294 | 32032 | S3 | Pre-coded security features. Only configurable parameters are Password Aging and automatic logout | 4 | LOW | 1 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_294 | BRN-LAS_294-S4-3-2 | #N/A | LAS_294 | 32032 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_294 | BRN-LAS_294-T2-2-6 | #N/A | LAS_294 | 32032 | T2 | Date and time can be changed | 2 | SIGNIFICANT | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_294 | BRN-LAS_294-T3-3-7 | #N/A | LAS_294 | 32032 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_294 | BRN-LAS_294-VD1-2-8 | #N/A | LAS_294 | 32032 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_294 | BRN-LAS_294-I1-1-5 | #N/A | LAS_294 | 32032 | I1 | Backup not accessible or not configured | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_299 | BRN-LAS_299-S1-1-1 | #N/A | LAS_299 | 34252 | S1 | Access to the system is not granted for the other user that operator. Administrator password is unknown | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_299 | BRN-LAS_299-S2-2-2 | #N/A | LAS_299 | 34252 | S2 | User management settings are not accessible or the most privileged access available at the time of the assessment has no access. | 2 | SIGNIFICANT | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_299 | BRN-LAS_299-S3-3-3 | #N/A | LAS_299 | 34252 | S3 | Security features are not accessible or the most privileged access available at the time of the assessment has no access. | 3 | MODERATE | 3 |  | Verify the system's security features compliance. If it's not, configure security features according to relevant SOP and good practices
BRN | BRN-LAS_299 | BRN-LAS_299-S4-2-4 | #N/A | LAS_299 | 34252 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_299 | BRN-LAS_299-T1-1-6 | #N/A | LAS_299 | 34252 | T1 | Audit Trail is not accessible or not configured within the system | 1 | HIGH | 6 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_299 | BRN-LAS_299-T3-2-7 | #N/A | LAS_299 | 34252 | T3 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 7 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_299 | BRN-LAS_299-VD1-1-8 | #N/A | LAS_299 | 34252 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_299 | BRN-LAS_299-I1-1-5 | #N/A | LAS_299 | 34252 | I1 | Server not assessed because not accessible. | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_299 | BRN-LAS_299-I2-1-6 | #N/A | LAS_299 | 34252 | I2 | Server not assessed because not accessible. | 1 | HIGH | 6 |  | Verify the status of the system's compliance. If necessary, change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_299 | BRN-LAS_299-I3-1-7 | #N/A | LAS_299 | 34252 | I3 | Server not assessed because not accessible. | 1 | HIGH | 7 |  | Verify the system's compliance status. If required, implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_299 | BRN-LAS_299-I4-1-8 | #N/A | LAS_299 | 34252 | I4 | Server not assessed because not accessible. | 1 | HIGH | 8 |  | Check the system's behaviour towards all electronic records. If autosave is not enforced by the system, document the manual workaround for saving output records
BRN | BRN-LAS_302 | BRN-LAS_302-S1-1-1 | #N/A | LAS_302 | 20161589 | S1 | Generic account present in the system | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_302 | BRN-LAS_302-S2-2-2 | #N/A | LAS_302 | 20161589 | S2 | Improper segregation of duties | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_302 | BRN-LAS_302-S3-3-3 | #N/A | LAS_302 | 20161589 | S3 | Server not assessed because not accessible. | 3 | MODERATE | 3 |  | Verify the system's security features compliance. If it's not, configure security features according to relevant SOP and good practices
BRN | BRN-LAS_302 | BRN-LAS_302-S4-2-4 | BRN-LAS_302-S4-2-1 | LAS_302 | 20161589 | S4 | Server not assessed because not accessible. | 2 | SIGNIFICANT | 4 |  | Verify the system's user List and ensure that is coherent and aligned with the users trained to use the system
BRN | BRN-LAS_302 | BRN-LAS_302-T1-1-9 | #N/A | LAS_302 | 20161589 | T1 | Server not assessed because not accessible. | 1 | HIGH | 9 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_302 | BRN-LAS_302-T2-1-10 | #N/A | LAS_302 | 20161589 | T2 | Server not assessed because not accessible. | 1 | HIGH | 10 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_302 | BRN-LAS_302-T3-2-11 | #N/A | LAS_302 | 20161589 | T3 | Server not assessed because not accessible. | 2 | SIGNIFICANT | 11 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_302 | BRN-LAS_302-VD1-1-12 | #N/A | LAS_302 | 20161589 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 12 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_302 | BRN-LAS_302-I1-1-5 | #N/A | LAS_302 | 20161589 | I1 | Backup is not accessible or the function is not configured. | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_302 | BRN-LAS_302-I3-1-6 | #N/A | LAS_302 | 20161589 | I3 | Data Folder is not protected from alteration/deletion within OS. | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_304 | BRN-LAS_304-S1-1-1 | #N/A | LAS_304 | 32503 | S1 | Generic login for SW application | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_304 | BRN-LAS_304-S2-2-2 | #N/A | LAS_304 | 32503 | S2 | 3 groups configured in the system, however the access to SW is generic so the proper segregation of duties cannot be assured | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_304 | BRN-LAS_304-S3-3-3 | #N/A | LAS_304 | 32503 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_304 | BRN-LAS_304-S4-2-4 | BRN-LAS_304-S4-3-1 | LAS_304 | 32503 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_304 | BRN-LAS_304-T1-1-7 | #N/A | LAS_304 | 32503 | T1 | Audit Trail is not accessible or not configured within the system | 1 | HIGH | 7 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_304 | BRN-LAS_304-T2-1-8 | #N/A | LAS_304 | 32503 | T2 | Time reference (Time zone) is not protected. Date and Time protected. | 1 | HIGH | 8 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_304 | BRN-LAS_304-T3-2-9 | BRN-LAS_304-T3-3-2 | LAS_304 | 32503 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_304 | BRN-LAS_304-VD1-1-10 | #N/A | LAS_304 | 32503 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_304 | BRN-LAS_304-I1-2-4 | #N/A | LAS_304 | 32503 | I1 | Backup is not accessible or the function is not configured. | 2 | SIGNIFICANT | 4 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_304 | BRN-LAS_304-I3-2-5 | #N/A | LAS_304 | 32503 | I3 | Data Folder is not protected from alteration/deletion within OS. | 2 | SIGNIFICANT | 5 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_305 | BRN-LAS_305-S2-3-1 | #N/A | LAS_305 | 33347 | S2 | User are assigned to at least 3 groups. Segregation of duties is not assured as some users have two roles assigned to its account | 3 | MODERATE | 1 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_305 | BRN-LAS_305-S3-4-2 | #N/A | LAS_305 | 33347 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 4 | LOW | 2 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_305 | BRN-LAS_305-S4-3-3 | BRN-LAS_305-S4-3-1 | LAS_305 | 33347 | S4 | No Official User List is in place | 3 | MODERATE | 3 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_305 | BRN-LAS_305-T2-2-6 | #N/A | LAS_305 | 33347 | T2 | Time reference (Time zone) is not protected. Date and Time protected. | 2 | SIGNIFICANT | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_305 | BRN-LAS_305-T3-3-7 | BRN-LAS_305-T3-3-2 | LAS_305 | 33347 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_305 | BRN-LAS_305-VD1-2-8 | #N/A | LAS_305 | 33347 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_305 | BRN-LAS_305-I1-1-5 | #N/A | LAS_305 | 33347 | I1 | The backup script is not protected from alteration by users not belonging to Admin group | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_305 | BRN-LAS_305-I3-1-6 | #N/A | LAS_305 | 33347 | I3 | Data not protected from alteration/deletion. Users (not Admin) can access to data folders and alter/delete data files. Moreover, also the backup script is not protected | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_323 | BRN-LAS_323-S1-1-1 | #N/A | LAS_323 | 30504 | S1 | Present active accounts with double profile | 1 | HIGH | 1 |  | If double accounts are necessary for business reasons, document and map their functionality and access capabilities, and grant traceability of their actions
BRN | BRN-LAS_323 | BRN-LAS_323-S2-2-2 | #N/A | LAS_323 | 30504 | S2 | No proper segregation of duties configured: '-User -Superuser (this group has the same privileges of Admin) -Admin -Minimal (nobody assigned) | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_323 | BRN-LAS_323-S3-3-3 | #N/A | LAS_323 | 30504 | S3 | Security features not configured properly: Password length: 6 digits Automatic log off: Not active Max attempts: 6 Password expiring time: 600 days | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_323 | BRN-LAS_323-S4-2-4 | BRN-LAS_323-S4-3-1 | LAS_323 | 30504 | S4 | The periodic review of authorized users is not performed properly | 2 | SIGNIFICANT | 4 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_323 | BRN-LAS_323-T2-1-7 | #N/A | LAS_323 | 30504 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_323 | BRN-LAS_323-T3-2-8 | BRN-LAS_323-T3-3-2 | LAS_323 | 30504 | T3 | Audit Trail is not routinely reviewed and there are no procedural requirements to execute the AT review | 2 | SIGNIFICANT | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_323 | BRN-LAS_323-VD1-1-9 | #N/A | LAS_323 | 30504 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_323 | BRN-LAS_323-I1-1-5 | #N/A | LAS_323 | 30504 | I1 | The backup script is not protected from alteration by users not belonging to Admin group | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_323 | BRN-LAS_323-I3-1-6 | #N/A | LAS_323 | 30504 | I3 | Data not protected from alteration/deletion. Users (not Admin) can access to data folders and alter/delete data files. Moreover, also the backup script is not protected | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_324 | BRN-LAS_324-S1-1-1 | #N/A | LAS_324 | 30501 | S1 | Present active accounts with double profile | 1 | HIGH | 1 |  | If double accounts are necessary for business reasons, document and map their functionality and access capabilities, and grant traceability of their actions
BRN | BRN-LAS_324 | BRN-LAS_324-S2-2-2 | #N/A | LAS_324 | 30501 | S2 | No proper segregation of duties configured: '-User -Superuser (this group has the same privileges of Admin) -Admin -Minimal (nobody assigned) | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_324 | BRN-LAS_324-S3-3-3 | #N/A | LAS_324 | 30501 | S3 | Security features not configured properly: Password length: 6 digits Automatic log off: Not active Max attempts: 6 Password expiring time: 600 days | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_324 | BRN-LAS_324-S4-2-4 | BRN-LAS_324-S4-3-1 | LAS_324 | 30501 | S4 | The periodic review of authorized users is not performed properly | 2 | SIGNIFICANT | 4 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_324 | BRN-LAS_324-T2-1-7 | #N/A | LAS_324 | 30501 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_324 | BRN-LAS_324-T3-2-8 | BRN-LAS_324-T3-3-2 | LAS_324 | 30501 | T3 | Audit Trail is not routinely reviewed and there are no procedural requirements to execute the AT review | 2 | SIGNIFICANT | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_324 | BRN-LAS_324-VD1-1-9 | #N/A | LAS_324 | 30501 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_324 | BRN-LAS_324-I1-1-5 | #N/A | LAS_324 | 30501 | I1 | The backup script is not protected from alteration by users not belonging to Admin group | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_324 | BRN-LAS_324-I3-1-6 | #N/A | LAS_324 | 30501 | I3 | Data not protected from alteration/deletion. Users (not Admin) can access to data folders and alter/delete data files. Moreover, also the backup script is not protected | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_325 | BRN-LAS_325-S1-1-1 | #N/A | LAS_325 | 30503 | S1 | Present active accounts with double profile | 1 | HIGH | 1 |  | If double accounts are necessary for business reasons, document and map their functionality and access capabilities, and grant traceability of their actions
BRN | BRN-LAS_325 | BRN-LAS_325-S2-2-2 | #N/A | LAS_325 | 30503 | S2 | No proper segregation of duties configured: '-User -Superuser (this group has the same privileges of Admin) -Admin -Minimal (nobody assigned) | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_325 | BRN-LAS_325-S3-3-3 | #N/A | LAS_325 | 30503 | S3 | Security features not configured properly: Password length: 6 digits Automatic log off: Not active Max attempts: 6 Password expiring time: 600 days | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_325 | BRN-LAS_325-S4-2-4 | BRN-LAS_325-S4-3-1 | LAS_325 | 30503 | S4 | The periodic review of authorized users is not performed properly | 2 | SIGNIFICANT | 4 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_325 | BRN-LAS_325-T2-1-7 | #N/A | LAS_325 | 30503 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_325 | BRN-LAS_325-T3-2-8 | BRN-LAS_325-T3-3-2 | LAS_325 | 30503 | T3 | Audit Trail is not routinely reviewed and there are no procedural requirements to execute the AT review | 2 | SIGNIFICANT | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_325 | BRN-LAS_325-VD1-1-9 | #N/A | LAS_325 | 30503 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_325 | BRN-LAS_325-I1-1-5 | #N/A | LAS_325 | 30503 | I1 | The backup script is not protected from alteration by users not belonging to Admin group | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_325 | BRN-LAS_325-I3-1-6 | #N/A | LAS_325 | 30503 | I3 | Data not protected from alteration/deletion. Users (not Admin) can access to data folders and alter/delete data files. Moreover, also the backup script is not protected | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_326 | BRN-LAS_326-S1-1-1 | #N/A | LAS_326 | 33271 | S1 | Present active accounts with double profile | 1 | HIGH | 1 |  | If double accounts are necessary for business reasons, document and map their functionality and access capabilities, and grant traceability of their actions
BRN | BRN-LAS_326 | BRN-LAS_326-S2-2-2 | #N/A | LAS_326 | 33271 | S2 | No proper segregation of duties configured: '-User -Superuser (this group has the same privileges of Admin) -Admin -Minimal (nobody assigned) | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_326 | BRN-LAS_326-S3-3-3 | #N/A | LAS_326 | 33271 | S3 | Security features not configured properly: Password length: 6 digits Automatic log off: Not active Max attempts: 6 Password expiring time: 600 days | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_326 | BRN-LAS_326-S4-2-4 | BRN-LAS_326-S4-3-1 | LAS_326 | 33271 | S4 | The periodic review of authorized users is not performed properly | 2 | SIGNIFICANT | 4 |  | Update the User List and ensure that is coherent and aligned with the users trained to use the system.
BRN | BRN-LAS_326 | BRN-LAS_326-T2-1-7 | #N/A | LAS_326 | 33271 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 7 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_326 | BRN-LAS_326-T3-2-8 | BRN-LAS_326-T3-3-2 | LAS_326 | 33271 | T3 | Audit Trail is not routinely reviewed and there are no procedural requirements to execute the AT review | 2 | SIGNIFICANT | 8 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_326 | BRN-LAS_326-VD1-1-9 | #N/A | LAS_326 | 33271 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_326 | BRN-LAS_326-I1-1-5 | #N/A | LAS_326 | 33271 | I1 | Information not provided | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_326 | BRN-LAS_326-I2-1-6 | #N/A | LAS_326 | 33271 | I2 | User can delete data | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_326 | BRN-LAS_326-I3-1-7 | #N/A | LAS_326 | 33271 | I3 | User can delete data | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_332 | BRN-LAS_332-S1-1-1 | #N/A | LAS_332 | 33724 | S1 | Local personal accounts to access to Windows. No user management to access to the SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_332 | BRN-LAS_332-S2-2-2 | #N/A | LAS_332 | 33724 | S2 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_332 | BRN-LAS_332-S3-3-3 | #N/A | LAS_332 | 33724 | S3 | Local personal accounts to access to Windows. No user management to access to the SW. | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_332 | BRN-LAS_332-S4-2-4 | #N/A | LAS_332 | 33724 | S4 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_332 | BRN-LAS_332-T1-2-8 | #N/A | LAS_332 | 33724 | T1 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 8 |  | Implement Audit Trail functionality within the SW or implement a paper based management in case of Audit Trail not fully compliant
BRN | BRN-LAS_332 | BRN-LAS_332-T2-1-9 | #N/A | LAS_332 | 33724 | T2 | Time Zone not protected | 1 | HIGH | 9 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_332 | BRN-LAS_332-T3-2-10 | #N/A | LAS_332 | 33724 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 10 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_332 | BRN-LAS_332-VD1-1-11 | #N/A | LAS_332 | 33724 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 11 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_332 | BRN-LAS_332-I1-1-5 | #N/A | LAS_332 | 33724 | I1 | Information not provided | 1 | HIGH | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_332 | BRN-LAS_332-I2-1-6 | #N/A | LAS_332 | 33724 | I2 | User can delete data | 1 | HIGH | 6 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_332 | BRN-LAS_332-I3-1-7 | #N/A | LAS_332 | 33724 | I3 | User can delete data | 1 | HIGH | 7 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_333 | BRN-LAS_333-S1-1-1 | #N/A | LAS_333 | 30974 | S1 | Local personal accounts to access to Windows. No user management to access to the SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_333 | BRN-LAS_333-S2-2-2 | #N/A | LAS_333 | 30974 | S2 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_333 | BRN-LAS_333-S3-3-3 | #N/A | LAS_333 | 30974 | S3 | Local personal accounts to access to Windows. No user management to access to the SW. | 3 | MODERATE | 3 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_333 | BRN-LAS_333-S4-2-4 | #N/A | LAS_333 | 30974 | S4 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_333 | BRN-LAS_333-T1-2-8 | #N/A | LAS_333 | 30974 | T1 | Local personal accounts to access to Windows. No user management to access to the SW. | 2 | SIGNIFICANT | 8 |  | Implement Audit Trail functionality within the SW or implement a paper based management in case of Audit Trail not fully compliant
BRN | BRN-LAS_333 | BRN-LAS_333-T2-1-9 | #N/A | LAS_333 | 30974 | T2 | Time Zone not protected | 1 | HIGH | 9 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_333 | BRN-LAS_333-T3-2-10 | #N/A | LAS_333 | 30974 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 10 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_333 | BRN-LAS_333-VD1-1-11 | #N/A | LAS_333 | 30974 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 11 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_333 | BRN-LAS_333-I2-1-5 | #N/A | LAS_333 | 30974 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 5 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_333 | BRN-LAS_333-I3-1-6 | #N/A | LAS_333 | 30974 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_333 | BRN-LAS_333-I4-1-7 | #N/A | LAS_333 | 30974 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 7 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_334 | BRN-LAS_334-S1-1-1 | #N/A | LAS_334 | 28845 | S1 | Generic account configured in the system. | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_334 | BRN-LAS_334-S2-2-2 | #N/A | LAS_334 | 28845 | S2 | Some users involved in the process are belonging also to Administrator group. Administrator+Supervisor | 2 | SIGNIFICANT | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_334 | BRN-LAS_334-S3-3-3 | #N/A | LAS_334 | 28845 | S3 | Password length: 6 digits Automatic log off: 20 minutes Max attempts: 5 Password expiring time: No expiration | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_334 | BRN-LAS_334-S4-2-4 | #N/A | LAS_334 | 28845 | S4 | Generic users on the system | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_334 | BRN-LAS_334-VD1-1-8 | #N/A | LAS_334 | 28845 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_334 | BRN-LAS_334-I2-1-5 | #N/A | LAS_334 | 28845 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 5 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_334 | BRN-LAS_334-I3-1-6 | #N/A | LAS_334 | 28845 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_334 | BRN-LAS_334-I4-1-7 | #N/A | LAS_334 | 28845 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 7 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_335 | BRN-LAS_335-S1-1-1 | #N/A | LAS_335 | 30396 | S1 | Generic account configured in the system. | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_335 | BRN-LAS_335-S2-2-2 | #N/A | LAS_335 | 30396 | S2 | Some users involved in the process are belonging also to Administrator group. Administrator+Supervisor | 2 | SIGNIFICANT | 2 |  | Update Utilization SOP defining User levels and related privileges. If multiple accounts for the same user are necessary for business flows, identify, map, and define them, making sure to grant traceability of their actions throughout the workflow.
BRN | BRN-LAS_335 | BRN-LAS_335-S3-3-3 | #N/A | LAS_335 | 30396 | S3 | Password length: 6 digits Automatic log off: 20 minutes Max attempts: 5 Password expiring time: No expiration | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_335 | BRN-LAS_335-S4-2-4 | #N/A | LAS_335 | 30396 | S4 | Generic users on the system | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_335 | BRN-LAS_335-VD1-1-8 | #N/A | LAS_335 | 30396 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_335 | BRN-LAS_335-I2-1-3 | #N/A | LAS_335 | 30396 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_335 | BRN-LAS_335-I3-1-4 | #N/A | LAS_335 | 30396 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_335 | BRN-LAS_335-I4-1-5 | #N/A | LAS_335 | 30396 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 5 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_336 | BRN-LAS_336-S3-3-1 | #N/A | LAS_336 | 31639 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 3 | MODERATE | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_336 | BRN-LAS_336-S4-2-2 | BRN-LAS_336-S4-4-1 | LAS_336 | 31639 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_336 | BRN-LAS_336-T2-1-6 | #N/A | LAS_336 | 31639 | T2 | Date and time can be changed | 1 | HIGH | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_336 | BRN-LAS_336-T3-2-7 | BRN-LAS_336-T3-4-2 | LAS_336 | 31639 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_336 | BRN-LAS_336-VD1-1-8 | #N/A | LAS_336 | 31639 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_336 | BRN-LAS_336-I2-1-3 | #N/A | LAS_336 | 31639 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_336 | BRN-LAS_336-I3-1-4 | #N/A | LAS_336 | 31639 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_336 | BRN-LAS_336-I4-1-5 | #N/A | LAS_336 | 31639 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 5 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_337 | BRN-LAS_337-S3-3-1 | #N/A | LAS_337 | 28847 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 3 | MODERATE | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_337 | BRN-LAS_337-S4-2-2 | BRN-LAS_337-S4-4-1 | LAS_337 | 28847 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_337 | BRN-LAS_337-T3-2-6 | BRN-LAS_337-T3-4-2 | LAS_337 | 28847 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 6 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_337 | BRN-LAS_337-VD1-1-7 | #N/A | LAS_337 | 28847 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 7 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_337 | BRN-LAS_337-I2-1-3 | #N/A | LAS_337 | 28847 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_337 | BRN-LAS_337-I3-1-4 | #N/A | LAS_337 | 28847 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_337 | BRN-LAS_337-I4-1-5 | #N/A | LAS_337 | 28847 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 5 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_338 | BRN-LAS_338-S3-3-1 | #N/A | LAS_338 | 33346 | S3 | Security features are not configured according to the relevant Security Policy (or equivalent document) | 3 | MODERATE | 1 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_338 | BRN-LAS_338-S4-2-2 | BRN-LAS_338-S4-4-1 | LAS_338 | 33346 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_338 | BRN-LAS_338-T2-1-6 | #N/A | LAS_338 | 33346 | T2 | Time reference (Time zone) is not protected. Date and Time protected. | 1 | HIGH | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_338 | BRN-LAS_338-T3-2-7 | BRN-LAS_338-T3-4-2 | LAS_338 | 33346 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_338 | BRN-LAS_338-VD1-1-8 | #N/A | LAS_338 | 33346 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_338 | BRN-LAS_338-I2-1-3 | #N/A | LAS_338 | 33346 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_338 | BRN-LAS_338-I3-1-4 | #N/A | LAS_338 | 33346 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_338 | BRN-LAS_338-I4-1-5 | #N/A | LAS_338 | 33346 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 5 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_339 | BRN-LAS_339-S1-1-1 | #N/A | LAS_339 | 20162373 | S1 | Present Generic account | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_339 | BRN-LAS_339-S2-2-2 | #N/A | LAS_339 | 20162373 | S2 | Present some users with Admin and developer privileges | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_339 | BRN-LAS_339-T2-1-6 | #N/A | LAS_339 | 20162373 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_339 | BRN-LAS_339-T3-2-7 | #N/A | LAS_339 | 20162373 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_339 | BRN-LAS_339-VD1-1-8 | #N/A | LAS_339 | 20162373 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_339 | BRN-LAS_339-I2-1-3 | #N/A | LAS_339 | 20162373 | I2 | Data is not protected from alteration/deletion by the SW application | 1 | HIGH | 3 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_339 | BRN-LAS_339-I3-1-4 | #N/A | LAS_339 | 20162373 | I3 | Data Folders are not protected from alteration/deletion within OS | 1 | HIGH | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_339 | BRN-LAS_339-I4-1-5 | #N/A | LAS_339 | 20162373 | I4 | The system force the user to save manually the file. It is not possible to close the SW without saving. However, if at the end of the analysis the data are not saved and the PC is turned off there is the risk to loose the data. | 1 | HIGH | 5 |  | Document the manual workaround for saving output records
BRN | BRN-LAS_340 | BRN-LAS_340-S1-1-1 | #N/A | LAS_340 | 20162374 | S1 | Present Generic account | 1 | HIGH | 1 |  | Remove generic accounts and configure unique and individual accounts for each user
BRN | BRN-LAS_340 | BRN-LAS_340-S2-2-2 | #N/A | LAS_340 | 20162374 | S2 | Present some users with Admin and developer privileges | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_340 | BRN-LAS_340-T2-1-6 | #N/A | LAS_340 | 20162374 | T2 | Time reference (Time zone) is not protected | 1 | HIGH | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_340 | BRN-LAS_340-T3-2-7 | #N/A | LAS_340 | 20162374 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_340 | BRN-LAS_340-VD1-1-8 | #N/A | LAS_340 | 20162374 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 8 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_340 | BRN-LAS_340-I1-2-5 | #N/A | LAS_340 | 20162374 | I1 | Backup settings not accessible or not configured | 2 | SIGNIFICANT | 5 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_340 | BRN-LAS_340-I3-2-6 | #N/A | LAS_340 | 20162374 | I3 | User has no access to OS level unless the keyboard is plugged-in via USB port. In OS level, user can alter the data | 2 | SIGNIFICANT | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_342 | BRN-LAS_342-S1-2-1 | #N/A | LAS_342 | 20000008 | S1 | Access to the system is not granted for the other user that operator. Administrator password is unknown | 2 | SIGNIFICANT | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_342 | BRN-LAS_342-S2-3-2 | #N/A | LAS_342 | 20000008 | S2 | User management settings are not accessible or the most privileged access available at the time of the assessment has no access. | 3 | MODERATE | 2 |  | Verify the system's user management configuration and its compliance. If it's not, configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_342 | BRN-LAS_342-S3-4-3 | #N/A | LAS_342 | 20000008 | S3 | Security features are not accessible or the most privileged access available at the time of the assessment has no access. | 4 | LOW | 3 |  | Verify the system's security features compliance. If it's not, configure security features according to relevant SOP and good practices
BRN | BRN-LAS_342 | BRN-LAS_342-S4-3-4 | BRN-LAS_342-S4-3-1 | LAS_342 | 20000008 | S4 | No Official User List is in place | 3 | MODERATE | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_342 | BRN-LAS_342-T1-2-7 | #N/A | LAS_342 | 20000008 | T1 | Audit Trail either not accessible or not configured. | 2 | SIGNIFICANT | 7 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_342 | BRN-LAS_342-T2-2-8 | #N/A | LAS_342 | 20000008 | T2 | Date and time can be changed using keyboard via USB port. | 2 | SIGNIFICANT | 8 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_342 | BRN-LAS_342-T3-3-9 | BRN-LAS_342-T3-3-2 | LAS_342 | 20000008 | T3 | Audit Trail either not accessible or not configured. | 3 | MODERATE | 9 |  | Verify if an Audit Trail Process is in place and if it's compliant. If not, implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_342 | BRN-LAS_342-VD1-2-10 | #N/A | LAS_342 | 20000008 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 10 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_342 | BRN-LAS_342-I1-2-3 | #N/A | LAS_342 | 20000008 | I1 | Backup is not configured | 2 | SIGNIFICANT | 3 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_342 | BRN-LAS_342-I3-2-4 | #N/A | LAS_342 | 20000008 | I3 | Data Folder is not protected from alteration/deletion within OS. | 2 | SIGNIFICANT | 4 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_348 | BRN-LAS_348-S3-4-1 | #N/A | LAS_348 | 32841 | S3 | System does not allow to configure any Security feature | 4 | LOW | 1 |  | Investigate with the supplier if it is possible to implement proper security features or issue a Security SOP regulating a manual or hybrid management of Password security management
BRN | BRN-LAS_348 | BRN-LAS_348-S4-3-2 | BRN-LAS_348-S4-2-1 | LAS_348 | 32841 | S4 | No Official User List is in place | 3 | MODERATE | 2 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_348 | BRN-LAS_348-T1-2-5 | #N/A | LAS_348 | 32841 | T1 | Audit Trail is not accessible or not configured within the system | 2 | SIGNIFICANT | 5 |  | Verify the system's Audit Trail functionality compliance. If it's not, implement Audit Trail functionality within the SW or implement a fully compliant paper based management in case of Audit Trail not compliant
BRN | BRN-LAS_348 | BRN-LAS_348-T2-2-6 | #N/A | LAS_348 | 32841 | T2 | Time reference is not protected (Time zone) | 2 | SIGNIFICANT | 6 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_348 | BRN-LAS_348-T3-3-7 | BRN-LAS_348-T3-2-2 | LAS_348 | 32841 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 7 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_348 | BRN-LAS_348-T4-3-8 | #N/A | LAS_348 | 32841 | T4 | Original Raw Data is not maintained in the SW. However, the data is being transferred by the lab operator to the NuGenesis platform | 3 | MODERATE | 8 |  | Implement a process of automatic data transfer following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual data transfer
BRN | BRN-LAS_348 | BRN-LAS_348-VD1-2-9 | #N/A | LAS_348 | 32841 | VD1 | Computer System Validation executed but incomplete | 2 | SIGNIFICANT | 9 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_348 | BRN-LAS_348-I1-1-2 | #N/A | LAS_348 | 32841 | I1 | Automatic backup not accessible or not configured. | 1 | HIGH | 2 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_348 | BRN-LAS_348-I3-1-3 | #N/A | LAS_348 | 32841 | I3 | Data can be modified within OS | 1 | HIGH | 3 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_354 | BRN-LAS_354-S4-2-1 | BRN-LAS_354-S4-3-1 | LAS_354 | 2002923 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 1 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_354 | BRN-LAS_354-T3-2-4 | #N/A | LAS_354 | 2002923 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 4 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_354 | BRN-LAS_354-VD1-1-5 | #N/A | LAS_354 | 2002923 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 5 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_354 | BRN-LAS_354-I1-1-2 | #N/A | LAS_354 | 2002923 | I1 | Automatic backup not accessible or not configured. | 1 | HIGH | 2 |  | Implement a process of automatic data backup following the relevant procedure and best practice. If not possible, formalize within an official document/procedure, the reason for the necessity of manual approach and define the procedure for a manual backup on a reliable support (e.g. cloud, server)
BRN | BRN-LAS_354 | BRN-LAS_354-I3-1-3 | #N/A | LAS_354 | 2002923 | I3 | Data can be modified within OS | 1 | HIGH | 3 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_355 | BRN-LAS_355-S4-2-1 | BRN-LAS_355-S4-3-1 | LAS_355 | 20050244 | S4 | No Official User List is in place | 2 | SIGNIFICANT | 1 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_355 | BRN-LAS_355-T3-2-4 | #N/A | LAS_355 | 20050244 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 2 | SIGNIFICANT | 4 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_355 | BRN-LAS_355-VD1-1-5 | #N/A | LAS_355 | 20050244 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 5 |  | Issue the missing documentation and perform the missing tests
BRN | BRN-LAS_355 | BRN-LAS_355-I2-1-5 | #N/A | LAS_355 | 20050244 | I2 | Data not protected | 1 | HIGH | 5 |  | Change permissions of users in order to prevent unauthorized actions from users and to prevent alteration/deletion of data, or formalize within an official document/procedure, using a risk-based approach, the rationale for the applied strategy
BRN | BRN-LAS_355 | BRN-LAS_355-I3-1-6 | #N/A | LAS_355 | 20050244 | I3 | Data folders can be deleted by all the users (generic Windows access to OS) | 1 | HIGH | 6 |  | Implement an automatic security policy for data protection on the server and/or on the local OS through active directories policies, OS folder restrictions or external security tool and configure User Privileges in such that users may not have the possibility to modify/delete data.
BRN | BRN-LAS_356 | BRN-LAS_356-S1-1-1 | #N/A | LAS_356 | 30705 | S1 | Generic Windows account to OS. No access control to SW. | 1 | HIGH | 1 |  | Update the application SW in order to implement a proper user management. If not applicable, issue a Security SOP regulating a manual or hybrid management of secondary security features including the management of a paper-based logbook
BRN | BRN-LAS_356 | BRN-LAS_356-S2-2-2 | #N/A | LAS_356 | 30705 | S2 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 2 |  | Configure at least 3 user groups with a proper segregation of duties
BRN | BRN-LAS_356 | BRN-LAS_356-S3-3-3 | #N/A | LAS_356 | 30705 | S3 | Generic Windows account to OS. No access control to SW. | 3 | MODERATE | 3 |  | Configure security features according to relevant SOP and good practices
BRN | BRN-LAS_356 | BRN-LAS_356-S4-2-4 | #N/A | LAS_356 | 30705 | S4 | Generic Windows account to OS. No access control to SW. | 2 | SIGNIFICANT | 4 |  | Issue a User List that is coherent and aligned with the users trained for the use of the system.
BRN | BRN-LAS_356 | BRN-LAS_356-T1-1-7 | #N/A | LAS_356 | 30705 | T1 | It has been not possible to check if the methods can have the same name or there is a versioning | 1 | HIGH | 7 |  | Investigate with the supplier if it is possible to implement a proper configuration of the Audit Trail or implement a paper based management in case of a not complete or not fully compliant Audit Trail
BRN | BRN-LAS_356 | BRN-LAS_356-T2-1-8 | #N/A | LAS_356 | 30705 | T2 | Time Zone not protected | 1 | HIGH | 8 |  | Configure Time Reference protection, e.g. locking Time reference setting and assigning modification privileges only to authorized users, or implement procedural mitigation.
BRN | BRN-LAS_356 | BRN-LAS_356-T3-3-9 | #N/A | LAS_356 | 30705 | T3 | Audit Trail is not routinely reviewed. Global procedure (DOC-000277399) not followed | 3 | MODERATE | 9 |  | Implement a fully compliant Audit Trail Review process
BRN | BRN-LAS_356 | BRN-LAS_356-VD1-1-10 | #N/A | LAS_356 | 30705 | VD1 | Computer System Validation executed but incomplete | 1 | HIGH | 10 |  | Issue the missing documentation and perform the missing tests

## Inventory (hidden)
LAS_001 | Broadmeadows | Chemistry
LAS_002 | Broadmeadows | Chemistry
LAS_003 | Broadmeadows | Chemistry
LAS_004 | Broadmeadows | BioChem
LAS_006 | Broadmeadows | BioChem
LAS_007 | Broadmeadows | BioChem
LAS_008 | Broadmeadows | BioChem
LAS_009 | Broadmeadows | Chemistry
LAS_010 | Broadmeadows | Chemistry
LAS_012 | Broadmeadows | BioChem
LAS_013 | Broadmeadows | Chemistry
LAS_014 | Broadmeadows | Chemistry
LAS_015 | Broadmeadows | Chemistry
LAS_016 | Broadmeadows | Chemistry
LAS_017 | Broadmeadows | Chemistry
LAS_018 | Broadmeadows | BioChem
LAS_019 | Broadmeadows | Chemistry
LAS_020 | Broadmeadows | Chemistry
LAS_021 | Broadmeadows | Chemistry
LAS_022 | Broadmeadows | Chemistry
LAS_023 | Broadmeadows | Chemistry
LAS_024 | Broadmeadows | BioChem
LAS_025 | Broadmeadows | BioChem
LAS_026 | Broadmeadows | Chemistry
LAS_027 | Broadmeadows | Chemistry
LAS_028 | Broadmeadows | BioChem
LAS_029 | Broadmeadows | BioChem
LAS_030 | Broadmeadows | BioChem
LAS_031 | Broadmeadows | BioChem
LAS_032 | Broadmeadows | BioChem
LAS_033 | Broadmeadows | BioChem
LAS_034 | Broadmeadows | BioChem
LAS_035 | Broadmeadows | BioChem
LAS_036 | Broadmeadows | Chemistry
LAS_038 | Broadmeadows | Chemistry
LAS_039 | Broadmeadows | Chemistry
LAS_040 | Broadmeadows | Chemistry
LAS_041 | Broadmeadows | Chemistry
LAS_042 | Broadmeadows | Chemistry
LAS_043 | Broadmeadows | Chemistry
LAS_046 | Broadmeadows | Microbiology
LAS_047 | Broadmeadows | BioChem
LAS_048 | Broadmeadows | BioChem
LAS_049 | Broadmeadows | BioChem
LAS_050 | Broadmeadows | BioChem
LAS_052 | Broadmeadows | Chemistry
LAS_055 | Broadmeadows | Chemistry
LAS_056 | Broadmeadows | Chemistry
LAS_057 | Broadmeadows | Chemistry
LAS_059 | Broadmeadows | BioChem
LAS_060 | Broadmeadows | BioChem
LAS_061 | Broadmeadows | Chemistry
LAS_062 | Broadmeadows | Chemistry
LAS_063 | Broadmeadows | Chemistry
LAS_064 | Broadmeadows | Chemistry
LAS_065 | Broadmeadows | Chemistry
LAS_067 | Broadmeadows | Chemistry
LAS_068 | Broadmeadows | Chemistry
LAS_069 | Broadmeadows | Chemistry
LAS_070 | Broadmeadows | Chemistry
LAS_071 | Broadmeadows | Chemistry
LAS_072 | Broadmeadows | Chemistry
LAS_073 | Broadmeadows | Chemistry
LAS_074 | Broadmeadows | Chemistry
LAS_075 | Broadmeadows | Chemistry
LAS_076 | Broadmeadows | Chemistry
LAS_077 | Broadmeadows | BioChem
LAS_078 | Broadmeadows | BioChem
LAS_079 | Broadmeadows | BioChem
LAS_080 | Broadmeadows | BioChem
LAS_081 | Broadmeadows | BioChem
LAS_082 | Broadmeadows | BioChem
LAS_083 | Broadmeadows | Microbiology
LAS_084 | Broadmeadows | Microbiology
LAS_085 | Broadmeadows | Microbiology
LAS_086 | Broadmeadows | BioChem
LAS_087 | Broadmeadows | Non-Production Area
LAS_089 | Broadmeadows | Chemistry
LAS_090 | Broadmeadows | Chemistry
LAS_091 | Broadmeadows | Chemistry
LAS_092 | Broadmeadows | Chemistry
LAS_093 | Broadmeadows | Non-Production Area
LAS_094 | Broadmeadows | Non-Production Area
LAS_095 | Broadmeadows | Non-Production Area
LAS_096 | Broadmeadows | Non-Production Area
LAS_097 | Broadmeadows | Non-Production Area
LAS_098 | Broadmeadows | Non-Production Area
LAS_099 | Broadmeadows | Non-Production Area
LAS_100 | Broadmeadows | Non-Production Area
PCS_001 | 999 PMCS (33A / 33B) | Toll Process
PCS_006 | 110 | Toll Process
PCS_007 | 110 | Toll Process
PCS_008 | 110 | Toll Process
PCS_009 | TOLL / PMCS | Toll Process
PCS_010 | 999 | Toll Process
PCS_011 | 999 | Toll Process
PCS_012 | 999 | Toll Process
PCS_013 | 999 PMCS | Toll Process
PCS_014 | 999 PMCS | Toll Process
PCS_015 | 999 PMCS | Toll Process
PCS_016 | 999 PMCS | Toll Process
PCS_017 | 999 PMCS | Toll Process
PCS_018 | 999 PMCS | Toll Process
PCS_019 | 999 PMCS | Toll Process
PCS_020 | 999 PMCS | Toll Process
PCS_021 | PFF and BMW | Toll Process
PCS_022 | 999 PMCS | BMF Process
PCS_023 | 999 PMCS | BMF Process
PCS_024 | 999 PMCS | BMF Process
PCS_025 | 240 | BMF Process
PCS_026 | 242 | BMF Process
PCS_027 | 287 | BMF Process
PCS_028 | 113 | BMF Process
PCS_029 | 113 | BMF Process
PCS_030 | 113 | BMF Process
PCS_031 | 113 | BMF Process
PCS_033 | 113 | BMF Process
PCS_034 | 113 | BMF Process
PCS_035 | 113 | BMF Process
PCS_036 | 113 | BMF Process
PCS_037 | 113 | BMF Process
PCS_039 | 113 | BMF Process
PCS_040 | 113 | BMF Process
PCS_041 | 113 | BMF Process
PCS_042 | 113 | BMF Process
PCS_043 | 113 | BMF Process
PCS_044 | 113 | BMF Process
PCS_045 | 113 | BMF Process
PCS_046 | 113 | BMF Process
PCS_047 | 113 | BMF Process
PCS_048 | 113 | BMF Process
PCS_049 | 113 | BMF Process
PCS_050 | BMF / PMCS | BMF Process
PCS_051 | 120 | Turner Facility
PCS_052 | 120 | Turner Facility
PCS_055 | Turner / PMCS | Turner Facility
PCS_056 | Turner / PMCS | Turner Facility
PCS_057 | 999 PMCS | Turner Facility
PCS_058 | 999 PMCS | Turner Facility
PCS_059 | 120 | Turner Facility
PCS_060 | Facility A / PMCS | Facility A (AlbuRx) Process
PCS_061 | 999 PMCS | Facility A (AlbuRx) Process
PCS_062 | 999 PMCS | Facility A (AlbuRx) Process
PCS_063 | 999 PMCS | Facility A (AlbuRx) Process
PCS_064 | 999 PMCS | Facility A (AlbuRx) Process
PCS_065 | 999 PMCS | Facility A (AlbuRx) Process
PCS_066 | 461 | Central Services
PCS_067 | 420 | Central Services
PCS_068 | 420 | Central Services
PCS_069 | 420 | Central Services
PCS_070 | 420 | Central Services
PCS_071 | 462 | Central Services
PCS_072 | Common / Utilities | Central Services
PCS_073 | PRIVIGEN Turner Facility | Central Services
PCS_074 | PRIVIGEN Turner Facility | Central Services
PCS_075 | Common / Utilities | Central Services (Tank Farm and Facility F)
PCS_079 | 650 / Utilities | Central Services (Tank Farm and Facility F)
PCS_080 | 650 / Utilities | Central Services (Tank Farm and Facility F)
PCS_081 | 650 / Utilities | Central Services (Tank Farm and Facility F)
PCS_082 | 651 / Utilities | Central Services (Tank Farm and Facility F)
PCS_083 | 651 / Utilities | Central Services (Tank Farm and Facility F)
PCS_084 | 654 / Utilities | Central Services (Tank Farm and Facility F)
PCS_085 | 203 | Filling and Packaging
PCS_086 | 203 | Filling and Packaging
PCS_087 | 203 | Filling and Packaging
PCS_088 | 203 | Filling and Packaging
PCS_089 | 203 | Filling and Packaging
PCS_090 | 203 | Filling and Packaging
PCS_091 | 203 | Filling and Packaging
PCS_092 | 203 | Filling and Packaging
PCS_093 | 203 | Filling and Packaging
PCS_094 | 203 | Filling and Packaging
PCS_095 | 170 | Filling and Packaging
PCS_096 | 170 | Filling and Packaging
PCS_097 | 170 | Filling and Packaging
PCS_098 | 190 | Filling and Packaging
PCS_099 | 203 | Filling and Packaging
PCS_102 | 102 (Dispensing) | Filling and Packaging
PCS_105 | 102 | Filling and Packaging
PCS_113 | 180 | Filling and Packaging
PCS_114 | 180 | Filling and Packaging
PCS_115 | 180 | Filling and Packaging
PCS_116 | 180 | Filling and Packaging
PCS_117 | 180 | Filling and Packaging
PCS_118 | 180 | Filling and Packaging
PCS_119 | 180 | Filling and Packaging
PCS_120 | 103 | Filling and Packaging
PCS_121 | 103 | Filling and Packaging
PCS_124 | 103 (Bottle Preparation) | Filling and Packaging
PCS_127 | 103 | Filling and Packaging
PCS_129 | TOLL / Packaging | Filling and Packaging
PCS_130 | TOLL / Packaging | Filling and Packaging
PCS_136 | 104 | Filling and Packaging
PCS_137 | 104 | Filling and Packaging
PCS_138 | 104 | Filling and Packaging
PCS_139 | 104 | Filling and Packaging
PCS_140 | 104 | Filling and Packaging
PCS_141 | TOLL / Packaging | Filling and Packaging
PCS_142 | TOLL / Packaging | Filling and Packaging
PCS_149 | 205 | Filling and Packaging
PCS_150 | 999 PMCS | Filling and Packaging
PCS_151 | 999 PMCS | Filling and Packaging
PCS_152 | Turner Facility/Facility A | Filling and Packaging
PCS_153 | 377 | Common Systems
PCS_154 | 377 | Common Systems
PCS_155 | 377 | Common Systems
PCS_156 | 377 | Common Systems
PCS_157 | 999 PMCS | Common Systems
PCS_158 | Common / PMCS | Common Systems
PCS_159 | Old Village | Common Systems
PCS_160 | New Village | Common Systems
PCS_161 | TOLL / PMCS | Common Systems
PCS_162 | Utilities, BMF / PMCS | Common Systems
PCS_163 | 301 | Common Systems
PCS_164 | 301 | Common Systems
PCS_165 | 302 | Common Systems
PCS_166 | 303 | Common Systems
PCS_167 | 304 | Common Systems
PCS_168 | 305 | Common Systems
PCS_169 | N/A | Common Systems
PCS_170 | 230 | Q-Fever Process
PCS_171 | 112 | Q-Fever Process
PCS_172 | 112 | Q-Fever Process
PCS_173 | 111 | Laboratories
PCS_175 | 999 / Facility F | Facility F (Base Fractionation Facility) Process
PCS_176 | 999 / Facility F | Facility F (Base Fractionation Facility) Process
PCS_177 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_178 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_179 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_180 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_181 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_182 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_183 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_184 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_185 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_186 | 641 / Pooling | Facility F (Base Fractionation Facility) Process
PCS_188 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_189 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_190 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_191 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_192 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_193 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_194 | 665 / General | Facility F (Base Fractionation Facility) Process
PCS_196 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_197 | 645 / Fractionation | Facility F (Base Fractionation Facility) Process
PCS_198 | 642 / DEAE (PC) Elution | Facility F (Base Fractionation Facility) Process
PCS_199 | 642 / DEAE (PC) Elution | Facility F (Base Fractionation Facility) Process
PCS_200 | 642 / DEAE (PC) Elution | Facility F (Base Fractionation Facility) Process
PCS_201 | Toll Facility | Filling line 6
PCS_202 | Toll Facility | Filling line 6
PCS_203 | Toll Facility | Filling line 6
PCS_204 | Toll Facility | Filling line 6
PCS_205 | Toll Facility | Autoclave 15
PCS_206 | Toll Facility | Motus trayoff unit
PCS_207 | Toll Facility | Toll/Filling line 6
PCS_208 | Toll Facility | Toll/Filling line 6
PCS_209 | Toll Facility | Toll/Filling line 6
PCS_210 | AlbuRX facility (A) | Filling
PCS_211 | AlbuRX facility (A) | Bulk
PCS_212 | AlbuRX facility (A) | Filling
PCS_213 | AlbuRX facility (A) | Bulk
PCS_214 | AlbuRX facility (A) | Bulk
PCS_215 | CSL BMW | Qualification/Requalification Department
PCS_216 | CSL BMW | Qualification/Requalification Department
PCS_217 | CSL BMW | Qualification/Requalification Department
PCS_218 | CSL BMW | Qualification/Requalification Department
PCS_219 | CSL BMW | Qualification/Requalification Department
PCS_220 | CSL BMW | Qualification/Requalification Department
PCS_221 | CSL BMW | Qualification/Requalification Department
PCS_222 | 203 | Filling and Packaging
PCS_223 | BMF Process | Filling and Packaging
PCS_224 | BMF Process | Filling and Packaging
PCS_225 | BMF Process | Filling and Packaging
PCS_242 | BMF Process | Filling and Packaging
PCS_243 | BMF Process | Filling and Packaging
PCS_244 | BMF Process | Filling and Packaging
PCS_245 | BMF Process | Filling and Packaging
PCS_246 | BMF Process | Filling and Packaging
PCS_247 | BMF Process | Filling and Packaging
PCS_248 | BMF Process | Filling and Packaging
PCS_249 | BMF Process | Filling and Packaging
PCS_250 | BMF Process | Filling and Packaging
PCS_251 | BMF Process | Filling and Packaging
PCS_252 | BMF Process | Filling and Packaging
PCS_253 | BMF Process | Filling and Packaging
PCS_254 | BMF Process | Filling and Packaging
PCS_255 | BMF Process | Filling and Packaging
PCS_256 | 642 / DEAE (PC) Elution | Facility F (Base Fractionation Facility) Process

## CSL Groups (hidden)
QC Chemistry
QC Biochemistry
QC Microbiology
QC Sample Management
QC Excellence
QC Analytical Technology
LES Lab Support
LES Supporting Systems
Fractionation
Execution Systems - Frac
Bulk
Execution Systems - Bulk
Biotechnology
Execution Systems - Biotechnology
Filling
Execution Systems - Filling
Visual Inspection
Execution Systems - Visual Inspection
Packaging
Execution Systems - Packaging
I & T
CSV